## Slide 1: CPAP Calculator Development
Data for Action Deep Dive
24th February 2026

## Slide 2: Objective & conceptual framework
Goal: Estimate the number of CPAP devices needed per facility to meet neonatal demand.
Approach: Combines data-driven metrics and expert opinion, stratified by facility neonatal care level (Level I*/II/III).
Uses: Facility-level admissions, patient weight categories, LOS, device functionality, surge factor, and neonatal level.
Designed for planning & resource allocation, not as exact counts for individual patients.

## Slide 3: Principles of the calculator
Weight-stratified eligibility:
  <2000g → criteria-based (clinical eligibility)
  ≥2000g → expert-opinion rates
Hybrid estimation: Higher of criteria/expert estimate vs. actual CPAP given.
Concurrent demand: Calculated via average daily usage × surge factor.
LOS-driven: Accounts for patient days on CPAP.
Maintenance buffer: Ensures functional availability (~3–5% device redundancy).
Protection mechanism
Eligible count is never lower than observed CPAP use. This avoids underestimating clinician-judged need.

## Slide 4: How concurrent demand is modeled
Model daily concurrency:
For each day:
  Identify all eligible babies admitted
  Count those still admitted (admission ≤ day ≤ discharge)
This naturally accounts for:
  Length of stay
  Overlapping admissions
  Congestion
We then compute:
  Mean daily concurrent use
  Peak daily concurrent use
  Surge ratio (peak / average)

## Slide 5: Equipment adjustment factors
Parameter | Source | Notes
Eligible (%) per weight category | High-performing facilities + expert | Median % used for formula
LOS (days) per weight category | High-performing facilities + expert | Median LOS applied
Surge factor | Operational data | Median peak / average daily usage, capped at 3
Maintenance buffer | Equipment data | Fraction of non-functional devices (~3–5%)
Monthly admissions | Facility-level data | Summed across weight categories
High-performing facilities (coverage >80%) are being used for calibration to reduce bias.

## Slide 6: Estimation formula
CPAP devices needed = (Monthly admissions × % eligible × LOSdays / 30) × Surge factor × (1 + Maintenance buffer)
Step 1: Determine eligible patients by weight category and neonatal level.
Step 2: Multiply by LOS → average daily concurrent patients.
Step 3: Apply surge factor → accounts for peaks; level-specific.
Step 4: Apply maintenance buffer → ensures functional devices available.

## Slide 7: Parameter data
Weight band | No. of facilities | % Eligible (median IQR) | 
<1000g | 20 | 100 (99-100) | 
1000-1499g | 11 | 100 (100-100) | 
1500-1999g | 63 | 26 (19-44) | 
2000-2499g | 51 | 16 (10-31) | 
2500-4000g | 38 | 10 (6-19) | 
>4000g | 22 | 20 (11-30) | 
Using top coverage performers (80-100%)
They reflect better CPAP device needs, picking from both CPAP eligible by analysis and CPAPs done by clinical judgement

## Slide 8: Parameter data
Level of neonatal care | % admissions eligible*
Level I | 17%
Level II | 15%
Level III | 22%
Using top coverage performers (80-100%)
Surge factor: 2.7
Buffer for device maintenance: +3.4% - from NEST360 devices data
Length of hospital stay: 5 days | Length of stay under CPAP?

## Slide 9: Scope, use & limitations
Level-dependent calibration: Most reliable for hospitals similar to high-performing facilities of same neonatal care level.
Case mix variability: Extreme deviations in weight distribution or acuity may reduce accuracy.
Surge and LOS assumptions: Median values used; actual peaks may differ by level.
Maintenance/functional rates: Facility-specific equipment may change needs.
Daily counts vs. formula: Formula estimates average/peak, not individual patient tracking.
Recommendation: Use as planning tool, not operational prescription; update parameters with local data, stratified by neonatal care level.

## Slide 10: Demonstration + comparison with other tools
Hospital | Admissions (2025) | ABC tool | NEST tool 2024 | NEST tool 2026 | Current CPAPs installed
304 - Mloganzila | 118 | 15-23 | 8 | 12 | 35
307 – Temeke | 246 | 15-38 | 13 | 17 | 14
323 - Ubungo | 32 | 5 | 4 | 3 | 5

## Slide 11: Discussion
Are the % eligible by levels of care reasonable from your experiences?
Could volume of admissions in either of the levels influence the % in 1 above? Especially level I and II?
Surge factor: Should it vary by facility care level?
Length of stay (LOS):
  thoughts on use of overall patient LOS (5 days) ?
  appreciate death as a competing risk
  under-treatment noted in Malawi (last year – time to CPAP analysis).

## Slide 12: Next steps
Improve the tool based on today’s discussion
Implement an excel based calculator
Start on other device categories with general learnings from CPAP and device category specific nuances.